ACCRÉDITATION SANS FRONTIÈRES · ISO 15189 READY

QC Rules (Westgard/Trend) and Actions

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | LAB-SOP-QC-01-EN |
| --- | --- |
| Title | QC Rules (Westgard/Trend) and Actions |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |

## Table of Contents

1. Purpose · 2. Scope · 3. Definitions · 4. Rule Application · 5. Trend and Shift Detection · 6. Action Record · 7. Review of Released Results · 8. Distribution of Responsibilities · 9. Accompanying Documentation

## 1. Purpose

To define the statistical rules applied to internal quality control results, what each rule violation means, and the actions required, so that every operator makes the same accept/reject decision for the same data.

## 2. Scope

All quantitative examinations with IQC under LAB-SOP-A-03. The rules below are the laboratory's default multi-rule set; a section may adopt a different rule set for a specific examination only with a documented justification approved by the Quality Manager.

## 3. Definitions

- 1-2s — one control result exceeds ±2 SD from target (warning rule)

- 1-3s — one control result exceeds ±3 SD (rejection: random error)

- 2-2s — two consecutive results (same level, or two levels in the same run) exceed the same ±2 SD limit (rejection: systematic error)

- R-4s — within one run, one level exceeds +2 SD and another exceeds −2 SD, a range of 4 SD (rejection: random error)

- 4-1s — four consecutive results exceed the same ±1 SD limit (rejection: systematic error or shift)

- 10-x — ten consecutive results fall on the same side of the target (rejection: shift or trend)

- Shift — a sudden and sustained change in the mean; Trend — a gradual drift in one direction over six or more results

## 4. Rule Application

| Rule | Status | Most likely cause | Action |
| --- | --- | --- | --- |
| 1-2s | Warning | Random variation — or the start of a problem | Do not reject alone. Inspect the other rules. Note on the chart |
| 1-3s | Reject | Random error: pipetting, bubble, clot, wrong control | Hold results. Repeat control with fresh material. If in limits, release; if not, investigate |
| 2-2s | Reject | Systematic error: calibration, reagent lot, temperature | Hold results. Recalibrate or change reagent. Repeat controls. Review released results since last good control |
| R-4s | Reject | Random error affecting one run | Hold results. Repeat both levels. Check for mechanical fault |
| 4-1s | Reject | Systematic error developing | Hold results. Recalibrate. Check reagent, lamp, temperature |
| 10-x | Reject | Shift or trend: reagent deterioration, drift, new control lot | Hold results. Investigate cause of shift. Re-establish target if control lot has changed |

## 5. Trend and Shift Detection

- Each section lead reviews Levey-Jennings charts weekly for patterns not yet triggering a rule: six or more results drifting in one direction, results clustering on one side of the target, increasing scatter

- A developing trend is investigated before it becomes a rejection: reagent expiry, calibrator lot, environmental temperature, analyser maintenance due

- A shift after a known event (new reagent lot, recalibration, service) is expected; the section lead confirms the new mean is within the allowable bias and re-establishes the target value if justified per LAB-SOP-A-03

## 6. Action Record

Every rule violation is recorded on LAB-SOP-QC-01-F01 with: date and time, analyser, examination, control level and lot, the rule violated, the investigation, the corrective action, the repeat control results, the decision, and the operator and reviewer. A rejection that is 'resolved' only by repeating the control until it passes, without identifying a cause, is not an acceptable record; if three repeats are needed the analyser is taken out of service pending investigation.

## 7. Review of Released Results

When a systematic error rule (2-2s, 4-1s, 10-x) is violated, patient results released since the last acceptable control run are listed. The section lead assesses, with a pathologist where needed, whether the error magnitude could have changed a clinical decision. Where it could, the requesting clinicians are informed, the specimens are re-examined if available, and corrected reports are issued per LAB-SOP-PO-04. The review is documented on the action record.

## 8. Distribution of Responsibilities

| Task | Quality Manager | Section Lead | Laboratory Staff |
| --- | --- | --- | --- |
| Applying rules at every control evaluation |  |  | Responsible |
| Investigating and documenting each rejection |  | Verify | Responsible |
| Weekly chart review for trends and shifts |  | Responsible |  |
| Reviewing released results after systematic error | Approve | Responsible |  |
| Approving any alternative rule set | Responsible | Propose |  |

## 9. Accompanying Documentation

| Code | Title |
| --- | --- |
| LAB-SOP-A-03 | Internal Quality Control (IQC) |
| LAB-SOP-QC-01-F01 | QC Rule Violation and Action Record |
| LAB-SOP-PO-04 | Corrected Reports |
| LAB-SOP-A-04 | External Quality Assessment |

FREE SAMPLE — This document is a complete, unrestricted sample from the ASF ISO 15189 READY documentation pack, shared so a facility can judge the depth and quality of the full pack before purchase. Replace every [BRACKETED] placeholder with your facility's own details. The full pack contains the complete document set, forms, master index, and clause crosswalk.