ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Instrument Reprocessing & Sterilization

Standard Operating Procedure / Policy

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | DENT-STR-01-EN |
| --- | --- |
| Title | Instrument Reprocessing & Sterilization |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |

## Table of Contents

1. Purpose · 2. Scope · 3. Definitions · 4. Point of Use and Transport · 5. Cleaning and Disinfection · 6. Packaging and Sterilisation · 7. Failed Cycle or Indicator · 8. Storage and Use · 9. Training, Validation and Audit · 10. Distribution of Responsibilities · 11. Accompanying Documentation

## 1. Purpose

To ensure every reusable dental instrument is cleaned, disinfected, inspected, packaged, sterilised, stored, and tracked by a validated process, so that no instrument used on a patient can transmit infection from a previous patient.

## 2. Scope

All reusable instruments, handpieces, and devices at every surgery. Single-use items are never reprocessed. Applies to every staff member who handles instruments at any stage.

## 3. Definitions

- Reprocessing — the full sequence from point-of-use to sterile storage

- Decontamination room — the dedicated room with a one-way dirty-to-clean flow; instruments never move backwards

- Class 5 (integrating) indicator — a chemical indicator placed inside every pack that responds to time, temperature, and steam together

- Biological indicator (BI) — a spore test that proves lethality of the cycle; run weekly and after any repair

- Load — one autoclave cycle; identified by date, cycle number, and autoclave

## 4. Point of Use and Transport

- Immediately after use, gross debris is wiped from instruments with a damp gauze; sharps are handled with forceps; instruments are placed in a closed, labelled, puncture-resistant container

- Instruments are kept moist (enzymatic pre-soak spray or foam) if cleaning will be delayed beyond 30 minutes

- Containers are carried to the decontamination room dirty zone without stopping; instruments are never left in a surgery sink

## 5. Cleaning and Disinfection

Cleaning removes the soil that would otherwise shield organisms from the steriliser. It is the most important step and the one most often shortened.

- Preferred: automated washer-disinfector, validated cycle, instruments opened and racked so all surfaces are exposed; cycle printout retained

- Alternative: ultrasonic cleaner with enzymatic detergent at the manufacturer's concentration and time, followed by rinse; the ultrasonic is foil-tested weekly and the solution changed at least daily

- Manual cleaning only where automated methods are unavailable: heavy-duty gloves, eye protection, long-handled brush, under the water surface to avoid splashing; never in a hand-hygiene sink

- Handpieces are cleaned and lubricated per the manufacturer's instructions, then sterilised; they are never wiped-and-reused

- After cleaning, every instrument is inspected under magnification and good light for residual soil, damage, or corrosion; any soiled instrument is re-cleaned; any damaged instrument is withdrawn

## 6. Packaging and Sterilisation

- Instruments are packaged in pouches or wraps validated for steam sterilisation, with a class 5 indicator inside every pack and an external process indicator on every pack; hinged instruments are open; packs are not overfilled

- Each pack is labelled with the autoclave, load number, date, and contents; labels are applied before the cycle

- Sterilisation is by a type B (vacuum) steam autoclave for wrapped and hollow items at 134°C for the validated hold time; a type N autoclave is used only for unwrapped solid instruments for immediate use

- Each load is recorded on DENT-STR-01-F01: date, cycle number, operator, contents, cycle printout attached, external indicator checked, and the operator's signature releasing the load

- Daily: a Bowie-Dick / helix test before the first load on a type B autoclave, recorded. Weekly: a biological indicator, incubated and recorded on DENT-STR-01-F02. After any repair, relocation, or failed test: three consecutive BI passes before use

## 7. Failed Cycle or Indicator

Any failed cycle, failed daily test, failed external indicator, or failed BI: the autoclave is taken out of service and labelled; every load since the last passed test is identified from the log, recalled from storage, and reprocessed; instruments already used on patients are identified and the Principal Dentist decides on patient notification and lookback under DENT-GOV-08; the cause is found and fixed; three consecutive BI passes are required before return to service. The event is recorded as a nonconformity.

## 8. Storage and Use

- Sterile packs are stored dry, in closed drawers or cabinets in the clean zone, away from sinks and splash, rotated first-in-first-out, and used within the validated shelf life (normally 12 months for pouches if intact)

- Before opening at the chairside, the clinician checks the pack is dry and intact and that the external indicator has changed; after opening, the class 5 indicator is checked before the instruments touch the patient — a pack whose internal indicator has not changed is not used and the load is investigated

- A wet, torn, or dropped pack is not sterile and is returned for reprocessing

## 9. Training, Validation and Audit

Every staff member who reprocesses instruments is trained on this SOP and assessed as competent by observed practice at induction and annually (DENT-HR-02-F01). The autoclave and washer-disinfector are validated on installation and re-validated annually by a competent engineer with certificates retained. The Practice Manager audits the logs weekly and observes the full process quarterly; results are reported to management review.

## 10. Distribution of Responsibilities

| Task | Principal Dentist | Practice Manager | Decontamination Lead | Clinician |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final | Prepare |  |  |
| Cleaning, inspection, packaging, sterilisation, logs |  |  | Responsible |  |
| Daily and weekly tests |  | Verify | Responsible |  |
| Failed-test recall and lookback decision | Responsible | Coordinate | Execute |  |
| Chairside pack checks before use |  |  |  | Responsible |
| Annual validation and quarterly observation |  | Responsible |  |  |

## 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| DENT-STR-01-F01 | Autoclave Load Record |
| DENT-STR-01-F02 | Biological Indicator Log |
| DENT-DUWL-01 | Dental Unit Waterline Infection Control |
| DENT-IPC-01 | Infection Prevention and Control Programme |
| DENT-GOV-08 | Nonconformity and Corrective Action |
| DENT-HR-02-F01 | Individual Training Record |

FREE SAMPLE — This document is a complete, unrestricted sample from the ASF ISO 9001 READY documentation pack, shared so a facility can judge the depth and quality of the full pack before purchase. Replace every [BRACKETED] placeholder with your facility's own details. The full pack contains the complete document set, forms, master index, and clause crosswalk.