ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Vaccine Cold Chain Management

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | PED-04-EN |
| --- | --- |
| Title | Vaccine Cold Chain Management |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of PED-04-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. Storage Temperature Ranges · 7. Monitoring

8. Temperature Excursion Response · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure every vaccine at [ORGANIZATION NAME] is stored, monitored, and handled within the correct temperature range at all times, consistent with WHO and CDC cold chain guidance [1–2].

# 2. Scope

This SOP applies to every vaccine and diluent stored, transported, or administered at [ORGANIZATION NAME], from receipt through administration.

# 3. Definitions and Abbreviations

•  Cold chain — the unbroken sequence of correct storage and transport conditions maintained from manufacture through administration [3]

•  Temperature excursion — any reading outside the required storage range, however brief

•  DDL — Digital Data Logger, a continuous-recording temperature monitoring device [4–5]

# 4. General Provisions

Both freezing and heat exposure permanently damage vaccines, and this damage leaves no visible sign — a compromised vaccine looks identical to a viable one [5]. This is precisely why continuous, documented monitoring is not optional: there is no way to visually confirm a vaccine's condition after a suspected excursion.

# 5. Distribution of Responsibilities

| Task | Director | Pediatric Lead | Nursing Staff | Quality Manager |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| DDL calibration and data review |  | Responsible |  |  |
| Daily temperature check/log |  |  | Responsible |  |
| Temperature excursion response |  | Responsible |  |  |
| Cold chain audit |  |  |  | Responsible |

# 6. Storage Temperature Ranges

| Storage Unit | Range | Notes |
| --- | --- | --- |
| Refrigerator | 2°C to 8°C (36°F to 46°F) | Thermostat set at the factory or midpoint setting to reduce excursion risk [1–2] |
| Freezer | –50°C to –15°C (–58°F to +5°F) | Applies only to vaccines specifically requiring freezer storage [1–2] |

A thermostat dial shows relative levels of coldness, not an actual temperature — the only way to know the real storage temperature is to measure it directly with a temperature monitoring device [1].

# 7. Monitoring

Manual temperature checks twice daily leave roughly 23 hours and 50 minutes of each day genuinely unmonitored — an excursion during that window can occur and resolve without anyone knowing a vaccine was compromised. Continuous digital data logger monitoring, not periodic manual checks alone, is the real standard [4–5].

Step 1.  Check and record the storage unit's minimum and maximum temperatures at the start of each workday

Step 2.  Download and review DDL data at least every two weeks, and immediately whenever an excursion is flagged

Step 3.  Retain DDL download records for a minimum of 3 years [4]

Step 4.  Conduct a vaccine inventory assessment at least monthly, and before placing any new vaccine order

# 8. Temperature Excursion Response

A vaccine involved in a suspected temperature excursion is never simply discarded, and never simply kept in use — it is labeled “DO NOT USE,” physically separated from other vaccines in the storage unit, and held until a formal viability determination is made [3].

Step 1.  Label the affected vaccines “DO NOT USE” immediately

Step 2.  Physically move them to a separate, clearly marked area of the storage unit, apart from unaffected stock

Step 3.  Document the excursion — duration, temperature reached, and storage unit affected

Step 4.  Contact the vaccine manufacturer or relevant immunization authority for a viability determination before any further use or disposal

# 9. Special Considerations

•  A cold chain monitor (CCM) included with a vaccine shipment is checked for any transit excursion on arrival, and discarded after checking — it is a one-time-use device, not reused [3]

•  Vaccine transport outside the clinic is avoided where possible, since every transport event increases real exposure risk to inappropriate conditions; where transport is necessary, it follows a defined emergency packing procedure

# 10. Process Effectiveness Indicators (KPI)

•  Daily temperature log completion rate (target: 100%)

•  DDL data review compliance (every two weeks)

•  Temperature excursions, tracked and each one closed with a documented viability determination

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| SOP-MED-01 | Medication Management |
| GOV-01 | Nonconformity & Corrective Action |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [PEDIATRIC LEAD NAME] | ___________ | ___________ |
| Agreed by: [QUALITY MANAGER NAME] | ___________ | ___________ |

References

[1] Centers for Disease Control and Prevention. Chapter 5: Vaccine Storage and Handling. Epidemiology and Prevention of Vaccine-Preventable Diseases (Pink Book).

[2] New York State Department of Health / CDC Vaccines for Children (VFC) Program. Vaccine Storage and Handling Requirements.

[3] Centers for Disease Control and Prevention. Chapter 5: Vaccine Storage and Handling — excursion labeling and quarantine protocol.

[4] CDC Vaccines for Children (VFC) Program. Digital Data Logger requirements — review frequency and retention period.

[5] SenseAnywhere. What Are the WHO Guidelines for Vaccine Cold Chain Temperature Monitoring? 2025.