ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Laboratory Internal Quality Control

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | SOP-LAB-02-EN |
| --- | --- |
| Title | Laboratory Internal Quality Control |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of SOP-LAB-02-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. Daily Internal Quality Control · 7. The Westgard Multirule

8. External Quality Assessment · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure every laboratory test result reported at [ORGANIZATION NAME] is genuinely reliable, through daily internal quality control and periodic external quality assessment, consistent with CLIA's individualized quality control framework and the real Westgard multirule statistical standard [1–4].

# 2. Scope

This SOP applies to every quantitative test performed at [ORGANIZATION NAME]'s in-house laboratory, across general chemistry, hematology, and any specialty testing offered on-site.

# 3. Definitions and Abbreviations

•  Levey-Jennings chart — the standard chart plotting control results against the established mean and standard deviation (SD) over time

•  Westgard multirule — a set of statistical decision rules, applied together, for deciding whether a run is in control or must be rejected [3–4]

# 4. General Provisions

Laboratory testing quality rests on three real, distinct safeguards: qualified personnel, internal quality control, and external quality assessment — no one of the three substitutes for the other two [1–2].

# 5. Distribution of Responsibilities

| Task | Director | Laboratory Lead | Lab Technician | Quality Manager |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| Westgard rule configuration |  | Responsible |  |  |
| Daily control testing |  |  | Responsible |  |
| EQA/PT enrollment and submission |  | Responsible |  |  |
| Out-of-control investigation |  | Responsible |  | Informed |

# 6. Daily Internal Quality Control

Two levels of control material are tested at least once each day of patient testing, before patient samples are run, for every non-waived test the laboratory performs, plotted on a Levey-Jennings chart against the established mean and SD [1–2].

# 7. The Westgard Multirule

[ORGANIZATION NAME]'s laboratory applies the real, widely-adopted Westgard multirule to decide whether a run is accepted or rejected, consistent with CLSI C24 guidance [3–4].

| Rule | What It Means | Detects |
| --- | --- | --- |
| 1-3s | One control value exceeds the mean ± 3 SD | Random error — run rejected [3–4] |
| 2-2s | Two consecutive control values exceed the same ±2 SD limit | Systematic shift — run rejected [3–4] |
| R-4s | The range between two control values exceeds 4 SD | Random error across levels — run rejected [3–4] |
| 4-1s | Four consecutive values exceed the same ±1 SD limit | Gradual systematic shift — run rejected [3–4] |

If any rejection rule is violated, patient results from that run are held — never released — the cause is investigated (calibration, reagent, control material, instrument maintenance), and the run is repeated only after the cause is identified and corrected [4].

# 8. External Quality Assessment

[ORGANIZATION NAME]'s laboratory enrolls in an approved external quality assessment or proficiency testing programme for every test category it performs. Unlike internal quality control, EQA uses blinded samples from an outside provider and compares results against peer laboratories — this is the only external, independent check on real testing accuracy [1–2]. Proficiency testing samples are tested by the same staff, methods, and frequency as routine patient specimens, never given special handling.

# 9. Special Considerations

•  Point-of-care testing devices used outside the main laboratory follow the same daily control principle, adapted to the device's specific quality control requirements

•  An unsatisfactory EQA/PT result triggers root cause investigation under the Nonconformity Procedure (GOV-01), not merely a repeat test

# 10. Process Effectiveness Indicators (KPI)

•  Daily QC completion rate (target: 100%)

•  Westgard rule violations, tracked and investigated

•  EQA/PT satisfactory result rate

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| SOP-LAB-01 | Specimen Collection & Chain of Custody |
| GOV-01 | Nonconformity & Corrective Action |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [LABORATORY LEAD NAME] | ___________ | ___________ |
| Agreed by: [QUALITY MANAGER NAME] | ___________ | ___________ |

References

[1] Centers for Medicare & Medicaid Services. Clinical Laboratory Improvement Amendments (CLIA) — Individualized Quality Control Plan. 42 CFR Part 493.

[2] International Organization for Standardization. ISO 15189:2022, Medical Laboratories — Requirements for Quality and Competence. Geneva: ISO; 2022.

[3] CASRAI. Westgard Rules: Multirule QC Explained. 2026.

[4] Westgard, J.O. et al. A Multi-Rule Shewhart Chart for Quality Control in Clinical Chemistry. Clinical Chemistry 27:493-501, 1981.