ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Clinical Microbiology Testing

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | SOP-LAB-12-EN |
| --- | --- |
| Title | Clinical Microbiology Testing |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of SOP-LAB-12-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. Gram Stain — The Real Preliminary Step · 7. Culture and Susceptibility Testing

8. Microbiology-Specific QC · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure microbiology testing at [ORGANIZATION NAME] — Gram stain, culture, identification, and antimicrobial susceptibility testing — is performed to a real standard that genuinely reflects the organism present, given that microbiology QC is fundamentally different from other laboratory sections [1–2].

# 2. Scope

This SOP applies to every microbiology specimen processed at [ORGANIZATION NAME], including blood cultures, wound cultures, urine cultures, and any other specimen submitted for organism identification and susceptibility testing.

# 3. Definitions and Abbreviations

•  AST — Antimicrobial Susceptibility Testing, determining which antibiotics are effective against an identified organism

•  Genotype-phenotype discordance — a real, documented phenomenon: molecular detection of a resistance gene does not always match actual phenotypic resistance behavior — phenotypic AST remains a critical component even when molecular methods are used, not a redundant step [3]

# 4. General Provisions

Microbiology QC is different in kind from chemistry or hematology QC: most microbiology testing is qualitative — whether a medium grows the correct organism, whether a stain reacts properly, whether a smear is read correctly — rather than a single measured number, and this shapes every element of this SOP [1].

A real, documented failure: a positive blood culture Gram stain was read as gram-negative rods, when the organism was actually gram-positive cocci — the stain's decolorizer reagent had degraded. No single person made an obvious error; a small, upstream QC step (decolorizer verification) had been skipped, and a patient was started on the wrong antibiotic. This is exactly what Section 8's QC requirements exist to prevent [1].

# 5. Distribution of Responsibilities

| Task | Director | Lab Lead | Technologist | - |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| Media and reagent QC oversight |  | Responsible |  |  |
| Gram stain reading and reporting |  |  | Responsible |  |
| Positive culture/critical result notification |  |  | Responsible |  |
| AST interpretation and reporting |  | Agreed | Responsible |  |

# 6. Gram Stain — The Real Preliminary Step

A Gram stain reported from a positive blood culture carries real, quantified accuracy — sensitivity around or above 90% even in a worst-case analysis — which is why it is trusted as an immediate preliminary result while full culture identification is pending, not merely a rough guess [4–5].

•  A positive blood culture and its Gram stain result are reported to the ordering clinician as a critical result, using the same closed-loop notification process as any other critical value — not a routine result released through standard channels [6]

•  A documented, common interpretive error is reporting only one organism when two or more are actually present — the reader continues searching the smear even after identifying one organism, since Gram-positive organisms are generally easier to visualize and can mask a co-present Gram-negative organism [7]

•  Gram stain of a specimen such as sputum can also assess whether the specimen itself is suitable for culture, based on the ratio of white cells to epithelial cells present [7]

# 7. Culture and Susceptibility Testing

| Stage | Real Timing |
| --- | --- |
| Gram stain (positive culture) | Reported immediately as a critical result [6] |
| Preliminary identification | Within 24 hours of positivity, per real, standard laboratory practice [8] |
| Full identification and AST | Within 48–72 hours from sample to final answer, using CLSI-approved methods and current CLSI M100 breakpoints [9–10] |
| Negative culture, final report | After 5 days of incubation with no growth [8] |

Where [ORGANIZATION NAME]'s microbiology work is sent to a reference laboratory rather than performed fully in-house, these timeframes govern the reference laboratory relationship under SOP-LAB-06, and the real 48–72 hour standard is used to assess whether turnaround is genuinely acceptable, not an arbitrary benchmark [9–10].

# 8. Microbiology-Specific QC

•  Culture media QC — each new lot tested against known reference strains before use, confirming the medium supports correct growth

•  Reagent and stain QC — including verification that Gram stain reagents, particularly the decolorizer, are functioning correctly — the specific, real failure mode described in Section 4 [1]

•  AST QC — tested against reference organisms with known susceptibility patterns, consistent with CLSI method standards [3, 9]

•  Personnel competency for smear reading and culture interpretation is addressed under SOP-LAB-11, including the molecular-specific technique focus where PCR-based identification methods are used

# 9. Special Considerations

•  Automated AST systems can show real variability depending on software version and card/panel used — results are interpreted with this in mind, not assumed uniformly reliable across all configurations [9]

•  A molecular resistance gene result and a phenotypic susceptibility result that disagree are both reported, with the discordance noted — phenotypic AST is not discarded in favor of a molecular result [3]

# 10. Process Effectiveness Indicators (KPI)

•  Gram stain to critical-result-notification time

•  Culture media and reagent QC completion, including decolorizer verification

•  Turnaround time from specimen receipt to final AST report, against the 48–72 hour benchmark

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| SOP-LAB-06 | Reference Laboratory & Send-Out Testing |
| SOP-LAB-10 | Laboratory Biosafety & Biohazardous Waste |
| SOP-LAB-11 | Laboratory Personnel Competency Assessment |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [LAB LEAD NAME] | ___________ | ___________ |
| Agreed by: [QUALITY MANAGER NAME] | ___________ | ___________ |

References

[1] Microbe Online. Quality Control in the Microbiology Laboratory: Concepts, Components, and How It All Fits — real decolorizer failure case, six QC components. 2026.

[2] Microbe Online. Quality Control in the Microbiology Laboratory — qualitative vs. quantitative QC distinction.

[3] PMC. Pooled Antibiotic Susceptibility Testing Performs Within CLSI Standards — genotype-phenotype discordance.

[4] Journal of Clinical Microbiology. First Notification of Positive Blood Cultures and the High Accuracy of the Gram Stain Report.

[5] PMC. First Notification of Positive Blood Cultures — sensitivity data.

[6] MLabs (University of Michigan). Blood Culture, Bacteria and Yeast — critical result classification, timing benchmarks.

[7] Journal of Laboratory Physicians / SSJM. Laboratory Perspective of Gram Staining and its Significance — multi-organism detection error, specimen quality assessment. 2014.

[8] MLabs (University of Michigan). Blood Culture, Bacteria and Yeast — preliminary result and negative culture timing.

[9] Wenzler et al. Antimicrobial susceptibility testing: An updated primer for clinicians — CLSI M02/M07, automated system variability. Pharmacotherapy, 2023.

[10] Rapid Phenotypic Antibiotic Susceptibility Profiling. PMC — CLSI-approved 48–72 hour AST turnaround benchmark.