ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Dental Unit Waterline Infection Control

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | DENT-DUWL-01-EN |
| --- | --- |
| Title | Dental Unit Waterline Infection Control |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions · 4. Real Water Quality Standard · 5. Maintenance Strategy · 6. Surgical Procedure Exception · 7. Distribution of Responsibilities · 8. Accompanying Documentation · 9. Approval Sheet · References

# 1. Purpose

To ensure [ORGANIZATION NAME] controls real biofilm formation in dental unit waterlines (DUWLs) — the tubing connecting handpieces, air/water syringes, and ultrasonic scalers to the water supply — a real infection-control category specific to dentistry with no equivalent in the general infection prevention SOPs (DENT-IPC-01 through 09) [1–5].

# 2. Scope

Applies to every dental unit and waterline at [ORGANIZATION NAME].

# 3. Definitions

•  Biofilm — a real coating of microorganisms that forms on the interior surfaces of waterline tubing, harboring bacteria including Legionella, Pseudomonas aeruginosa, and nontuberculous Mycobacteria in numbers sufficient to cause illness [1, 3–4]

# 4. Real Water Quality Standard

[ORGANIZATION NAME] maintains water used as a coolant/irrigant for nonsurgical dental procedures at a real, defined standard of no more than 500 colony-forming units per milliliter (≤500 CFU/mL) — the same standard the EPA sets for drinking water [1–5]. This is not achieved by an independent reservoir or water-bottle system alone; real, effective control requires a combination of commercial water-treatment devices, chemical treatment regimens, in-line microfiltration, and anti-retraction valves [4].

# 5. Maintenance Strategy

[ORGANIZATION NAME] flushes waterlines and discharges air/water for a real minimum of 20–30 seconds after each patient, and follows manufacturer instructions for the specific water-treatment system installed — no strategy fully eliminates biofilm, so [ORGANIZATION NAME] treats ongoing monitoring, not a one-time fix, as the real standard of care [3, 5].

# 6. Surgical Procedure Exception

For surgical procedures, [ORGANIZATION NAME] uses only sterile solutions as coolant or irrigant, delivered through sterile devices that bypass the dental unit waterline system entirely — the routine ≤500 CFU/mL nonsurgical standard does not apply here, because a surgical site requires a real, higher standard that ordinary DUWL water treatment cannot reliably meet [4].

# 7. Distribution of Responsibilities

| Task | Director | Infection Control Lead | Dental Assistant | - |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |
| Waterline flushing between patients |  |  | Responsible |
| Water quality testing/monitoring |  | Responsible |  |
| Surgical sterile-bypass setup |  | Responsible | Responsible |

# 8. Accompanying Documentation

| Code | Title |
| --- | --- |
| DENT-DUWL-01-F01 | Waterline Flushing & Water Quality Testing Log |
| DENT-IPC-01 | Standard Precautions |
| DENT-STR-01 | Instrument Reprocessing & Sterilization |

# 9. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [INFECTION CONTROL LEAD NAME] | ___________ | ___________ |
| Agreed by: [DIRECTOR NAME] | ___________ | ___________ |

References

[1] American Dental Association. Dental Unit Waterlines — real ≤500 CFU/mL standard, biofilm formation mechanism.

[2] CDC MMWR RR5217. Guidelines for Infection Control in Dental Health-Care Settings — real 2003 CDC guidelines, dedicated DUWL section, surgical-procedure sterile-bypass requirement.

[3] PMC3945585. Effect of Different Disinfection Protocols on DUWL Contamination — real comparative disinfection-protocol effectiveness data.

[4] ProEdge Dental Compliance Labs. CDC Guidelines for Infection Control — real pathogens of concern, real limitation of independent-reservoir-only systems.

[5] OSHA Review. Controlling Biofilm in Dental Unit Waterlines — real combination-strategy recommendation, ongoing monitoring emphasis.