ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Point-of-Care Testing Governance

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | H-LAB-05-EN |
| --- | --- |
| Title | Point-of-Care Testing Governance |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of H-LAB-05-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. The POCT Coordinator · 7. Competency — Not Auto-Recertification

8. Quality Control · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure every point-of-care test performed at [ORGANIZATION NAME] — outside the core laboratory, by clinical staff — meets the same real standard of accuracy and competence as a core-laboratory result, consistent with ISO 15189:2022's dedicated POCT requirements [1–2].

# 2. Scope

This SOP applies to every point-of-care device used at [ORGANIZATION NAME] — glucose meters, urine dipsticks, rapid antigen/antibody tests, pregnancy tests, and any other test performed at the bedside rather than in the core laboratory.

# 3. Definitions and Abbreviations

•  POCT — point-of-care testing; any test performed at or near the patient, outside the core laboratory, typically by clinical staff who are not laboratory specialists [1–3]

•  Auto-recertification — a real, documented unreliable practice: an operator self-declares ongoing competency without direct observation — explicitly distinguished from genuine competency assessment in this SOP [4–5]

# 4. General Provisions

POCT is real laboratory testing performed outside the core lab, and ISO 15189:2022 treats it as integral to the laboratory's quality system, not a peripheral convenience — governance, quality assurance, and training are explicitly required, and this is a documented, common area of finding during accreditation assessment [1–3].

Auto-recertification is a real, documented unreliable practice: a real study found that although every operator observed had been auto-recertified, only 20% actually met the 80% compliance threshold when directly, independently observed [4]. [ORGANIZATION NAME] does not rely on self-declared competency.

# 5. Distribution of Responsibilities

| Task | Director | POCT Coordinator | Lab Lead | Device Operators |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| POCT training program and competency checklist |  | Responsible | Agreed |  |
| Device QC and maintenance oversight |  | Responsible |  |  |
| Day-to-day testing and daily QC |  |  |  | Responsible |
| Reporting atypical/out-of-range results |  |  | Responsible | Responsible |

# 6. The POCT Coordinator

[ORGANIZATION NAME] appoints a dedicated POCT Coordinator — a real, evidence-supported role, not an administrative formality: a real study found sites with a dedicated coordinator achieved meaningfully higher operator compliance (82%) than sites without one (67%) [5]. The coordinator oversees the POCT program across every device type and every department using them, reporting into the core laboratory's quality system, not operating as a separate, parallel structure.

# 7. Competency — Not Auto-Recertification

Step 1.  Initial training covers device operation, why (not only how) QC is performed, routine maintenance and troubleshooting, and correct result documentation [6]

Step 2.  Competency is assessed through direct, independent observation using a structured checklist — never through operator self-declaration alone [4–5]

Step 3.  Competency is reassessed at a defined interval — every 6–12 months, and more frequently for low-volume POCT devices where skill decay is a real, documented risk [6]

Step 4.  A documented record of each competency assessment is retained per the record's defined retention period

# 8. Quality Control

•  Daily QC is performed and documented before patient testing begins on any POCT device — most modern devices are designed to refuse operation until QC has passed, and this built-in control is never defeated or bypassed [7]

•  QC records are reviewed for accuracy, not merely completion — a real, documented finding shows QC compliance can appear high (95%+) while the underlying records are actually inaccurate [8]

•  An atypical or out-of-range result triggers defined troubleshooting before being acted upon clinically, not immediate clinical action on a potentially erroneous number [8]

# 9. Special Considerations

•  Reagent and test strip management — lot tracking, expiry monitoring, storage conditions — is a genuinely common real failure point; a real study found roughly half of departments monitored had insufficient management of these materials [8]

•  Operator-level differences in technique produce real, measurable differences in result accuracy, even on identical devices — competency assessment is not a formality but addresses a documented, real source of error [6]

# 10. Process Effectiveness Indicators (KPI)

•  Operator competency assessment completion rate via direct observation (target: 100%, not auto-recertification)

•  Daily QC completion and record accuracy, audited periodically, not assumed from completion alone

•  Atypical result rate, tracked and reviewed for pattern by device and operator

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| H-LAB-02 | Laboratory Internal QC |
| H-HR-01 | Physician Credentialing |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [POCT COORDINATOR NAME] | ___________ | ___________ |
| Agreed by: [LAB LEAD NAME] | ___________ | ___________ |

References

[1] SimplerQMS. ISO 15189: Definition, Requirements, Implementation, and Accreditation — POCT governance requirements. 2026.

[2] Biomedical Scientist. Point-of-care testing and the new ISO 15189:2022 — Annex A, common area of finding. 2026.

[3] ISO-Library. ISO 15189:2022 — Medical Laboratories — Quality and Competence Requirements — POCT oversight.

[4] Assia, S. et al. Glucose point-of-care meter operators competency: An assessment checklist. PMC, 2020.

[5] Assia, S. et al. Glucose point-of-care meter operators competency — POCT coordinator compliance comparison. ScienceDirect, 2020.

[6] American Diabetes Association. Clinical Need for Point-of-Care Testing for Diabetes in CLIA-Waived Settings — training components, recertification interval. 2025.

[7] PMC. Point-of-Care Blood Glucose Testing for Diabetes Care in Hospitalized Patients — device QC lockout design.

[8] Lim, et al. What We Should Consider in Point of Care Blood Glucose Test — reagent management and QC record accuracy findings. PMC, 2021.