ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Laboratory Personnel Competency Assessment

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | H-LAB-11-EN |
| --- | --- |
| Title | Laboratory Personnel Competency Assessment |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of H-LAB-11-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. The Six Real Elements · 7. Schedule

8. Molecular Testing — Additional Focus · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure every person performing laboratory testing at [ORGANIZATION NAME] is demonstrated, through direct observation and real evidence, to actually perform each test accurately — not merely trained on it — consistent with the six real competency elements defined in CLIA [1–2].

# 2. Scope

This SOP applies to every person performing non-waived laboratory testing at [ORGANIZATION NAME], for every specific test system they are approved to perform — distinct from HR-01's general credentialing, which verifies qualifications on hire.

# 3. Definitions and Abbreviations

•  Training vs. education vs. competency assessment — three genuinely distinct concepts: training teaches how to perform a task; education provides theoretical background; competency assessment validates, through direct evidence, that the task is actually performed accurately — training alone does not satisfy this requirement [3]

# 4. General Provisions

Competency assessment is one of the most commonly cited deficiencies in real CAP and TJC laboratory inspections — incomplete assessments, improper timing, and unqualified assessors are the recurring, documented causes [4]. All six elements must genuinely be addressed for every test each person performs — partial completion does not satisfy the requirement [5].

# 5. Distribution of Responsibilities

| Task | Director | Lab Lead | Qualified Assessor | Testing Personnel |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| Assessment schedule maintenance |  | Responsible |  |  |
| Conducting the six-element assessment |  |  | Responsible |  |
| Undergoing assessment for every approved test |  |  |  | Responsible |
| Documentation review |  | Responsible |  |  |

# 6. The Six Real Elements

All six elements are genuinely addressed — not a documentation checkbox, but real, verifiable evidence for each [1–2].

| # | Element | Real Requirement |
| --- | --- | --- |
| 1 | Direct observation | The assessor physically watches the person handle real patient specimens — virtual observation does not satisfy this element; record review alone cannot substitute [6] |
| 2 | Result recording/reporting | How results, including critical results, are documented, verified, and communicated to clinical staff |
| 3 | Record review | Intermediate test results, proficiency testing results, QC records, and preventive maintenance records |
| 4 | Instrument maintenance observation | Direct observation of the person performing maintenance and function checks, not merely asking if they know how |
| 5 | Blind/proficiency testing | Test performance assessed via previously analyzed specimens, internal blind samples, or external proficiency testing |
| 6 | Problem-solving evaluation | Written tests, instrument troubleshooting, or sample-handling issue resolution |

# 7. Schedule

•  A new hire is assessed at 6 months and again at 12 months during their first year — semi-annual, not the standard annual cycle [4–5]

•  After the first year, assessment is annual for every test the person is approved to perform

•  Assessments are scheduled by discipline across defined months of the year, avoiding the real, documented problem of a single employee facing multiple simultaneous department competencies [5]

# 8. Molecular Testing — Additional Focus

Molecular/PCR testing is genuinely manual and technique-dependent, making personnel error a real, significant source of result variance — competency observation for molecular testing specifically covers pipetting technique to prevent aerosol contamination, adherence to unidirectional clean-to-dirty workflow, and correct biological safety cabinet use, not merely whether the correct result was produced [7].

# 9. Special Considerations

•  Competency assessment is not restricted to department supervisors alone — any qualified person may serve as an assessor, and limiting the pool unnecessarily reduces real assessment capacity [4]

•  A person returning after an extended absence from patient testing (approximately 12 months or more) is treated as a new hire for competency purposes, restarting the semi-annual cycle — time away is a real, documented gap in demonstrated proficiency, not a presumption of continued skill [8]

# 10. Process Effectiveness Indicators (KPI)

•  Competency assessments completed within their required timeframe (target: 100%)

•  All six elements documented for every completed assessment, not a partial subset

•  Assessor pool size, tracked to avoid over-reliance on a single individual

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| H-HR-01 | Physician Credentialing |
| H-LAB-05 | Point-of-Care Testing Governance |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [LAB LEAD NAME] | ___________ | ___________ |
| Agreed by: [QUALITY MANAGER NAME] | ___________ | ___________ |

References

[1] Today's Clinical Lab. Timing of Competency Assessments — 42 CFR §493.1413/1451, six elements. 2020.

[2] Illinois Department of Public Health. What Do I Need to Do to Assess Personnel Competency? — full six-element definition.

[3] Molecular MLS ASCP. Competency — training vs. education vs. competency assessment distinction.

[4] Laboratory Management Consultants. Common COLA and CAP Inspection Deficiencies — assessor pool, schedule requirements. 2026.

[5] Laboratory Management Consultants. Laboratory Staff Competency Assessment — direct observation requirement, no virtual substitute. 2026.

[6] Laboratory Management Consultants. Laboratory Staff Competency Assessment — element definitions.

[7] Molecular MLS ASCP. Competency — molecular-specific observation focus: pipetting, workflow, biosafety cabinet use.

[8] Lab Symplified. A CAP Inspector's Guide to Competency Assessment — returning-employee treatment as new hire.