ACCRÉDITATION SANS FRONTIÈRES · ISO 9001 READY

Method Validation & Verification

Standard Operating Procedure

Approved by Director, [ORGANIZATION NAME] [DIRECTOR NAME] ______________________ Date: __________

| Document Code | H-LAB-15-EN |
| --- | --- |
| Title | Method Validation & Verification |
| Edition | 1st Edition |
| Copy # | ___ |
| Effective Date | [EFFECTIVE DATE] |
| Next Review Date | [NEXT REVIEW DATE] |
| Organization | [ORGANIZATION NAME], Georgia |
| Language Pair | Full parallel version of H-LAB-15-KA — same content, not a summary |

# Table of Contents

1. Purpose · 2. Scope · 3. Definitions and Abbreviations

4. General Provisions · 5. Distribution of Responsibilities

6. What Requires Validation — No Exceptions · 7. Validation Parameters

8. Approval and Retention · 9. Special Considerations

10. Process Effectiveness Indicators (KPI) · 11. Accompanying Documentation

12. Approval Sheet · References

# 1. Purpose

To ensure every nonwaived test, method, or instrument at [ORGANIZATION NAME] is validated or verified before use in patient testing, with no exception based on how long the method has been available or previously used elsewhere [1].

# 2. Scope

This SOP applies to every nonwaived test, method, or instrument system used at [ORGANIZATION NAME], including a new instrument of the same make and model as one already in use, and any temporary replacement (loaner) instrument.

# 3. Definitions and Abbreviations

•  Validation — establishing performance characteristics for a method not previously validated by an outside body (e.g., a laboratory-developed test)

•  Verification — confirming that a method already validated by its manufacturer (e.g., an FDA-cleared kit) performs as claimed in this specific laboratory, on this specific equipment [1–2]

# 4. General Provisions

There is no exception for validation or verification of a test based on when it was first introduced by the laboratory — this real requirement applies regardless of how long the method has been in use, including an instrument of the same make and model already validated elsewhere, and including a temporary replacement (loaner) instrument [1].

# 5. Distribution of Responsibilities

| Task | Director | Medical Director | Lab Lead | Technologist |
| --- | --- | --- | --- | --- |
| Approval of this SOP | Final |  |  |  |
| Validation/verification plan design |  | Agreed | Responsible |  |
| Conducting the validation study |  |  |  | Responsible |
| Final report review and approval |  | Responsible |  |  |

# 6. What Requires Validation — No Exceptions

•  A new laboratory-developed test (LDT)

•  A new FDA-cleared kit or method being verified for use in this laboratory

•  A new instrument, including one of the same make and model as an existing instrument [1]

•  A temporary replacement (loaner) instrument — the same real requirement applies, not a reduced or waived process [1]

This requirement does not apply to traditional methods such as manual microscopy, conventional microbiologic cultures, conventional cytogenetics, and manual hematology and immunology tests — a real, specific, limited exclusion, not a general exemption for older methods [1].

# 7. Validation Parameters

A validation or verification study addresses the applicable parameters for the method [2–3]:

Precision (repeatability and reproducibility)

Accuracy (agreement with a known reference or comparator method)

Analytical sensitivity and specificity, and where clinically applicable, clinical sensitivity and specificity

Reportable range and linearity

Reference intervals, verified or established for [ORGANIZATION NAME]'s own patient population

# 8. Approval and Retention

Step 1.  An umbrella validation SOP defines the general process, customized to each specific test being validated [4]

Step 2.  Results are compiled into a validation report

Step 3.  The medical director reviews and approves the report before the test is added to the active test menu (H-LAB-14) [4]

Step 4.  Validation records are retained for as long as the method is in use, and for at least 2 years after it is discontinued — not merely until the next review cycle [1]

# 9. Special Considerations

•  Waived testing follows the manufacturer's instructions as written and does not require the same validation study, but is not exempt from competency and QC requirements under other SOPs [1]

•  If a laboratory-developed test is used, the real US regulatory disclaimer language is included on the report where applicable, noting the test was developed and its performance characteristics determined by [ORGANIZATION NAME], has not been FDA cleared or approved, and is used for clinical purposes under CLIA high-complexity authorization [1]

# 10. Process Effectiveness Indicators (KPI)

•  Every active test menu item with a current, retrievable validation report (target: 100%)

•  Validation completed before a new instrument or loaner enters patient testing, with zero exceptions

# 11. Accompanying Documentation

| Code | Title |
| --- | --- |
| H-LAB-14 | Test Menu & Scope of Testing |
| H-LAB-07 | Laboratory Equipment Calibration & Maintenance |

# 12. Approval Sheet

| Role | Date | Signature |
| --- | --- | --- |
| Prepared by: [LAB LEAD NAME] | ___________ | ___________ |
| Agreed by: [MEDICAL DIRECTOR NAME] | ___________ | ___________ |

References

[1] CAP. Test Method Validation and Verification — no-exception requirement, retention period, traditional-method exclusion.

[2] USPTO. Systems and methods for laboratory assay validation or verification — CAP/CLIA/ISO 15189 scope.

[3] Accountable HQ. Clinical Laboratory Configuration Management — validation parameters.

[4] Tear Based Sample Collection (clinicaltrials.gov protocol). Validation process for an LDT — umbrella SOP, medical director approval, LDT disclaimer.